Libra-1+

Single-phase Watt hour Meter

HANDBOOK

Wasion Technology (Shenzhen) Co., Ltd.

The information contained in this document is subject to change without notice.

Wasion makes no warranty of any kind with regard to this material, including, but not limited to, the implied warranties of merchantability and fitness for a particular purpose. Wasion will not be liable for errors contained herein or incidental consequential damages in connection with the furnishing, performance, or use of this material.

This document contains proprietary information that is protected by copyright. All rights are reserved. No part of this document may be photocopied or otherwise reproduced without consent of Wasion

Document Edition

| index | date | Remarks |
| --- | --- | --- |
| 1 | 2009-11-30 | First edition |
| 2 | 2009-12-20 | Adjust format |

Copyright© 2007 by Wasion

Published in a limited copyright sense, and all rights, including trade secrets are reserved.

Table of Contents

1.0 Product Description	5

1.1 General Information	5

1.1.1	ANSI Standard Compatible	7

1.1.2	Product List	8

1.1.3	Physical Description	8

1.1.4	Meter Forms	10

1.1.5	Basic Data Sheet	10

1.2 Normal Operating Mode	11

1.3 Software Tool	12

1.4 Technical Information	12

1.4.1 Theory of Operation	12

2 Operating Instructions	15

2.1 Mechanical Structure	15

2.1.1 Cover	15

2.1.2 AMR module, Relay(Optional)	17

2.2 Nameplate Information and Labels	17

2.2.1 Nameplate Information	17

2.3 Display	18

2.3.1 Display elements	18

2.3.2 Energy Accumulation	19

3.0 Maintenance Instructions	20

3.1 Test Procedures	20

3.1.1 Meter Testing Tools	20

3.2 Field Accuracy Test	21

3.3 Shop Test	21

3.3.1 Meter Shop Equipment	21

3.3.1.1 Equipment Setup	22

3.3.1.2 Testing	22

3.3.2 General Watt-hour Meter Test Procedure	22

3.4 Service	23

3.5 Repair	23

3.6 Returning a Meter	23

3.7 Cleaning	24

3.8 Storage	24

3.9 Troubleshooting Guide	24

3.9.1 Self-Tests and Error Codes	25

4.0 Installation Instructions	25

4.1 General	25

4.2 Inspection	25

4.3 Storage	26

4.4 Unpacking	26

4.5 Selecting a site	26

4.6 Meter installation	26

5.0 Site Analysis Guides and Diagram	27

5.1 Form 1S (Self-Contained)	28

5.2 Form 2S (Self-Contained)	28

5.3 Form 3S (Transformer Rated)	29

5.4 Form 3S (Transformer Rated)	29

5.5 Form 4S (Transformer Rated)	30

5.6 Diagrams	31

6.0 Appendix	32

# 1.0 Product Description

The LIBRA meter is a solid-state, single-phase residential electricity meter that provides utilities with unparalleled digital accuracy, reliability, serviceability and cost-effectiveness.

The LIBRA meter is one of the most adaptable meters on the residential market, providing an array of communications, auto-disconnect and application options to meet current and future business requirements.

The LIBRA metering performance meets and exceeds ANSI industry standards (and ANSI protocol) with 50 or 60 Hz models available to operate at 120vac and 240vac for Class 20,

Class 100, Class 200 and Class 320 socket based residential or commercial installations.

It is designed to be compatible with new or existing socket based single-phase residential or commercial metering installations. The advanced electronic design offers superior accuracy and performance advantages over electromechanical meter.

Figure 1-1: Libra Electronic Single-phase Watt hour Meter

## 1.1 General Information

The LIBRA meter was engineered specifically for the widely varied needs of the residential market. The LIBRA-I, the basic platform for all LIBRA meters, is a solid state electricity meter that measures single-phase energy consumption.

LIBRA meter is a breakthrough-metering platform designed to grow and change with the utility industry. With its low starting watts and low watts lost, the LIBRA meter captures more energy consumption than was measured in the past by electromechanical meters

- Reliability - Unique simplified sensor design and mechanical construction; no welded joints in terminal blades

- Low starting watts - Captures energy consumption at levels not registered by electromechanical meters

- Low burden - Minimizes utility system losses

- Tamper-resistant design - Minimizes theft of service. Programmable to register energy when meter inverted (turning the meter upside down)

- Friendly display –Large LCD, two line display, more easy to read

- Nocturnal visibility - When the environment low light, the LCD backlight automatically open.

- Remote Disconnect (optional, 200A max) - In many markets, and in future expected markets, it is necessary to be able to interrupt the flow of power through the meter, disconnecting the home or business from the grid. Additionally, the Libra meter can be programmed to automatically disconnect power. The disconnect switch is fully integrated inside the Libra-1+. This is a factory-installed option that must be specified at the time of order. To take advantage of all of the functionality this option offers, a two-way AMR device and system should be employed.

- Network Applications - The Libra is available in 12S and 25S forms for network applications.

- Four options for Energy Accumulation - Delivered only, Delivered + Received,

Delivered – Received, Received only); one option must be specified at time of order, but changes can be made after the fact via Wasion MeterBuddy® Software.

Boosts operational efficiency - Factory programming minimizes handling

- Inventory management - Compatible with existing LIBRA-I+® cover stock .the meter measures kWh/kVarh energy and is rated as an ANSI C12.10 class 0.5 meter. However, it is expected that accuracy under many operating conditions and loads will typically be within 0.2%. The meter is configured at the factory. Resetting energy is accomplished with MeterBuddy® software (Version 1.0 or greater) and some keys which used to operate meter. The Liquid Crystal Display (LCD) indicates energy consumption,current rate,used days,meter state.The display is covered in detail in Chapter 2, Operation Instructions. The 6 large characters of the LCD display alphanumeric information. A configuration port is located at the 7 o’clock position of the meter face. The factory uses the configuration port to configure the meter for a specific application. It is also used in conjunction with a reset adaptor and MeterBuddy® software to reset accumulated energy to zero

### ANSI Standard Compatible

Meets applicable standards:

- ANSI C12.1-1995

- ANSI C12.10-1997

- ANSI C12.18-1996

- ANSI C12.19-1997

- ANSI C12.20(Class 0.5)-1998

- ANSI C37.90.1-1989

- ANSI C62.45-1992

- IEC 61000-4-4

- IEC61000-4-2

- FCC Part 15,Subclass C

- DL/T645-2007

- DL/T698-1999

### Product List

The LIBRA meter is available with a series of options, offering the ideal solution for a range of different residential needs. All kind of product list show as the following tables

| TYPE | Class | Volts | Wire | Form | AMR Port | Disconnect |
| --- | --- | --- | --- | --- | --- | --- |
| Libra-1a | 100 | 120/240 | 2 | 1S | No | No |
| Libra-1+ Libra-1+V3.0 Libra-1b Libra-1c | 200 | 240 | 3 | 2S | Optional | Optional |
| Libra-1b Libra-1c | 320 | 240 | 3 | 2S | Optional | No |
| Libra-1a | 20 | 120/240 | 2 | 3S | No | No |
| Libra-1b Libra-1c | 20 | 120/240 | 3 | 4S | Optional | Optional |
| Libra-1c | 200 | 120 | 3 | 12S | Optional | Optional |
| Libra-1c | 200 | 120 | 3 | 25S | Optional | Optional |

### Physical Description

The meter uses a poly carbonate cover. The cover is molded in one piece.

The meter base assembly contains the CT and a relay. The dimensions show as below.

|  | A | B | C | D | E | F | G |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Libra-1+ Libra-1+V3.0 | 6.29” | 6.95” | 3.93” | 4.93” | 5.93 | 6.29” | 6.95” |
|  | 16cm | 17.7cm | 10cm | 12.54cm | 15.08cm | 16cm | 17.7cm |
| Libra-1a Libra-1b Libra-1c | 6.29” | 6.95” | 1.54” | 3.23” | 4.23” | 6.29” | 6.95” |
|  | 16cm | 17.7cm | 5.67cm | 8.21cm | 10.75cm | 16cm | 17.7cm |

### Meter Forms

The meter is available in several variants, which can be selected depending upon the application. The ANSI Standard S Base Meter Forms are shown and the Non-ANSI

Meter Forms without test links are shown in Table below

### Basic Data Sheet

The basic physical description of the meter and available S base forms have been described above.

General Specifications:

Active energy / Reactive energy meter

Received kWh / Delivered kWh / Delivered kWh + Received kWh / Delivered kWh -

Received kWh meter

Digital Multiplication Measurement Technique

Non-Volatile Memory

Designed for 15+ years life

AMR Platform:

AMR platform (Modular and Integrated)

| Item | Specifications |
| --- | --- |
| Power Requirements | Rating Voltage: 240V/220V/120V Frequency: 60Hz/50Hz Operating Voltage: ±20% Operating Range: ±3Hz |
| Test Amps | - 50 Amps for Class 320 - 30 Amps for Class 200 - 15 Amps for Class 100 - 2.5 Amps for Class 20 It is printed on the nameplate and referred to as “TA” |
| Operating Environment | Temperature: -20°to +75℃ Humidity: 0% to 95% non-condensing |
| Accuracy | ANSI C12.20 - 0.5 Accuracy Class |
| General LCD Display | six-digit Two lines Liquid Crystal Display Data Digit Height: 9.67mm |
| Characteristic Data | Starting Watts: 5 Watts rating current:200 A Starting Watts: 2.5 Watts rating current:100 A |
| Temperature Rise Specifications | Meets ANSI C12.1 Section 4.7.2.9 |
| Burden Data (apples to FM2S 240V/220V Meter) | Voltage Circuit: Watts: <0.5 Voltage: 240V/220V VA:<5 Current Coil-Self Contained Test Amp Current: 60Hz Service: 3-Wire Test Current (Amps): 30 VA: < 0.50 |
| Weight | 0.9Kg |

## 1.2 Normal Operating Mode

The meter accumulates and displays kWh. The meter provides various features, some of which are configurable at the factory.

- kWh energy measurement accumulate in 4 ways:

1. Received kWh: The Received energy as a positive quantity (In ANSI C12.10 standard “Delivered” means from consumer to utility)

2. Delivered kWh: The Delivered energy (In ANSI C12.10 standard “Received” means from utility to consumer)

3. Delivered + Received kWh: The Imported and exported energy are added together.

4. Delivered – Received kWh: The net energy is calculated

- Display of Energy is 6 integrals.

- Alarm, Disconnect indicators on LED Lamp

- Reset energy to zero in field – The customer can reset the energy accumulation to zero on site with MeterBuddy® Software.

- Simple energy overflow – The energy will restart from 000,000 again.

- IR, which used to communicate with the MeterBuddy® software.

- PLC(Power line Communication), which used to communicate with the server.

## 1.3 Software Tool

The meter is supported by the MeterBuddy® Software (version 1.0 or greater), which facilitates the resetting of the energy accumulation to zero, calibrate and other functions. Refer to Reading and Programming Instruction manual for MeterBuddy®.

There is a reset adaptor that is available, part number WH109-1. The adaptor fits over a optical port to connect to PC’s RS232 port.

## 1.4 Technical Information

### 1.4.1 Theory of Operation

The theory of operation of Single Phase meter is explained in conjunction with the

block diagram shown in Figure 1-3.

Figure 1-3 Meter Block Diagram

#### 1.4.1.1 Sensing Devices

The current transformer feed the scaled current signals to the meter chip senses current. The voltage signal is scaled for measurement using a precision resistor divider circuit

#### 1.4.1.2 Meter Chip

- Single Converter Technologywith a 21-bit delta-sigma ADC

- 4 analog inputs plus auxiliary inputs

- On-chip temperature measurement for digital temperature compensation

- Precision voltage reference

- 32-bit compute engine

- 8051-based MPU (80515), 1 clock cycle per instruction

- Real Time Clock (RTC)

- 152 segments for LCD drive

- UARTs.

- Voltage comparator(s)

- Hardware Watch Dog Timer.

- Synchronous serial interface (SSI) for real time data.

- Real-time monitoring of CE variables

#### 1.4.1.3 EEPROM

The ANSI C12.10 meter is equipped with a non-volatile memory that is used to store

the metered data, calibration parameters, configuration constants and the program parameters. The non-volatile memory does not require a battery to retain information

when line power is not present.

Schmitt Trigger, Filtered Inputs for Noise Suppression

Write Protect Pin for Hardware and Software Data Protection

High Reliability

Data Retention: 40 Years

#### 1.4.1.4 Relay

The Libra-1+V3.0 meter is equipped with a 200A(max) relay that is used to remote-disconnect the power supply

#### 1.4.1.5 LCD

LCD is designed to meet customer needs e.g. The balance of electricity, Cumulative total electricity consumption ,Multi-Rate electricity consumption,But the ANSI C12.10 meter only implement basic function.

#### 1.4.1.6 Infrared optical port

The optical port meets to C12.18-1996 physical layer standard. This port is used to communication with computer tools. The computer tools can be used to program meter.

#### 1.4.1.7 AMR Uart Port

The AMR port is a UART, which meets to DL/T645-2007 standard. This port can be used to communication with inner AMR module, such as RF/PLC/Zigbee etc.

#### 1.4.1.8 Pulses output LED

The Libra meter is equipped a IR that is used to kWh pulse output.

# 2 Operating Instructions

## 2.1 Mechanical Structure

### 2.1.1 Cover

See Figure 2-1 for a graphic representation of the meter cover.

Figure 2-1 Meter Cover

Description:

1) Photo Sensor, nocturnal visibility, When the environment low light, the LCD backlight automatically shine.

2) Non-touch Key, press your hand on the cover to trigger this key.

A. Libra-1a/b/c meter: same as Display switch. See as below 3).

B. Libra-1+ meter: When the status of Connect / Disconnect relay is “Armed”, consumer an use this key to confirm the action, which will close the relay.(This is a non-touch switch)

Ｃ.lLibra-1+v3.0 meter. Reverse

3) Display Switch: Reverse.

4) IR, Optical communication with PC software or Handle (hand-held) read device. and, it is used to output the kWh pulse.

In default mode, this port is always output the kWh pulse. When the meter gets the

optical communication signal, it will switch to communication mode automatically.

5) Energy Reset Key, Used to clear energy.

6) Test Key, switch to Test Mode.

7) Alarm/Disconnect LED Indicator, LED On indicators relay disconnect, LED Blinking indicators Relay connect and Balance is Alarming, LED off when Relay connect and Balance isn’t Alarming.

### 2.1.2 AMR module, Relay(Optional)

Libra can equipped with all kinds of AMR modules, and the relay is from 60Amax to

200Amax.

## 2.2 Nameplate Information and Labels

### 2.2.1 Nameplate Information

See Figure 2-2 for a graphic representation of the meter nameplate. The meter

nameplate is found on the front of the meter.

Figure 2-2 Meter Nameplate

The nameplate information is shown in Figure 2-2. The following numbered list coincides with the numbers in the figure.

- CL200 - Current Class

- 240V - Nominal Voltage operating range

- 3W - Number of wires for the metered service

- TYPE LIBRA - The type of meter: Libra

- TA30 - Test amperes:30A

- Kh2.0 – Watt-hour test constant

- CA 0.5 - ANSI C12.20 Accuracy Class S-base CA 0.5

- FM2S - ANSI C12.10 Form Number

- 60Hz - Nominal Frequency

## 2.3 Display

### 2.3.1 Display elements

The Liquid Crystal Display (LCD) indicates energy consumption, instantaneous

power, demand, frequency and other. All values are refreshed every second.

Main Display

Non-Peak Tariff

Peak Tariff

Menu ID

Menu ID

Register Success

Peso

PLC failure

Comm-Indicator

Battery Low

Units

Main Display

Supply power days

Description:

Menu ID: The default Display ID and it’s contents are listed as below.

| ID | Display Format | Content |
| --- | --- | --- |
| 00 | Erxxxx | Error Information |
| 01 | xx. xx. xx | Year. Month. Date |
| 02 | ××:××:×× | Hour: Minute: Second |
| 03 | ×××××× | Total 12 digits Meter Number，Two-screen display |
| 04 | ×××××× |  |
| 05 | ×××× | Kh value, unit: imp/kWh |
| 06 | ××××××kWh | 6 integrals, electric quantity unit |
| 07 | ××××××kWh | Cumulative Peak rate electricity consumption |
| 08 | ××××××kWh | Cumulative Off-Peak rate electricity consumption |
| 09 | ××××.×× P | The balance of electricity |
| 10 | ××××××kWh | Previous thirty(30) day-period peak consumption |
| 11 | ××××××kWh | Previous thirty(30) day-period Off-peak consumption |
| 12 | ××××××kWh | Current thirty(30) day-period peak consumption |
| 13 | ××××××kWh | Current thirty(30) day-period Off-peak consumption |
| 14 | Xx Days | Number of days into the current month-period |

### 2.3.2 Energy Accumulation

Cumulative total electricity consumption: Delivered + Received. The magnitudes of the Received energy and Delivered energy are added together. This mode of operation is often referred to as “unidirectional” because the meter registration will always be positive, regardless of the direction of energy flow. This accumulation mode may deter tampering by meter inversion.

Cumulative Peak rate electricity consumption: Peak period rates energy is accumulated individually.

Cumulative Off-peak rate electricity consumption: Off-Peak period rates energy is accumulated individually.

Current thirty(30) day-period peak Consumption: Current thirty(30) day-period peak period rates energy is accumulated individually.

Current thirty(30) day-period Off-peak consumption: Current thirty(30) day-period Off-peak period rates energy is accumulated individually

Last month peak, off-peak tariff consumption:  On setting dump days and billing time , meter dump the current peak/off-peak consumption to the last month peak/off-peak register, and then reset the current peak/off-peak register .if power off at dump day,  meter only dump the once after power on, regardless of whether power off duration cross dump days

Last thirty day consumption load profile: meter can record 30 day energy consumption load profile(interval:1 day),  .if power off at 24:00 daily,  meter only save the once after power on, regardless of  power off duration

# 3.0 Maintenance Instructions

WARNING: The information contained within this document is intended to be an aid to Qualified metering personnel. It is not intended to replace the extensive training necessary to install or remove meters from service. Any work on or near energized meters, meter sockets, or other metering equipment presents the danger or electrical shock. All work on these products must be performed by qualified industrial electricians and metering specialists only.

All work must be done in accordance with local utility safety practices and the procedures

outlined in the current edition of the Handbook for Electricity Metering.

## 3.1 Test Procedures

The Procedures described below are suggested test procedures for use with the Libra meter. They are not intended to replace local utility operating practices.

### 3.1.1 Meter Testing Tools

The meter is equipped with a light-emitting diode (IR LED) for verifying calibration

#### 3.1.1.1 Calibration LED

The Infra-Red LED emits optical pulses proportional to kWh accumulation. The output of

the calibration LED is compatible with commercially available test equipment designed

to verify the calibration of watt-hour meters in the shop and field. Each calibration pulse

represents Kh watt-hours of energy accumulated by the meter. The duration of each output pulse is approximately 10 milliseconds

## 3.2 Field Accuracy Test

The most accurate way to verify meter calibration is the use of test equipment designed

to compare the meter accumulation to a precision watt-hour standard. This equipment

uses the calibration pulse output from the meter to gate (start and stop) the standard at

precise times to minimize the measurement uncertainty. The equipment will also

typically provide known regulated loads to the meter and standard during the test. In the

event such equipment is not available, the Libra meter provides two methods for making

a reasonable verification of the meter’s accuracy.

1. Disk analog testing

2. Instantaneous demand feature of the Libra.

## 3.3 Shop Test

Shop testing consists of verifying the meter’s accuracy.

### 3.3.1 Meter Shop Equipment

The Libra meter has a capacitive type power supply designed to operate at a single nominal supply voltage. The low burden design of the Libra should be compatible with

most commercially available test equipment. Meters may be tested in any shop that meets the requirements outlined in the current editions of the Handbook for Electricity Metering published by the Edison Electric Institute and the American National Standard Code for Electricity Metering.

### 3.3.1.1 Equipment Setup

The meter mounting equipment and its electrical connections must be used as required for the meter form number on the meter nameplate. For some test equipment, the meter

test link(s) must be opened in order to isolate the meter voltage and current circuits.

Refer to the instructions provided by the test equipment manufacturer to determine if

this is necessary. The Libra meter is also available in non-standard forms without test links where improved reliability and/or tamper resistance is desired.

### 3.3.1.2 Testing

The Watt-hour constant (Kh) of a meter is defined as Watt-hours per disk revolution.

Because electronic meters do not rely on disk revolutions to measure energy, Kh is rinted on the meter label as a reference to an equivalent electromechanical meter as required by applicable meter standards.

### 3.3.2 General Watt-hour Meter Test Procedure

To test the meter, proceed as follows:

1. Note the meter Kh value listed on the nameplate.

2. Select the desired voltage and current level(s) on the test equipment. Observe the appropriate voltage and current ratings for the Device Under Test. Exceeding the device ratings can result in permanent damage to the meter.

3. Install the meter in the test socket, making certain that the socket is wired and/or configured for the appropriate meter form.

4. Align the optical pickup of the test equipment with the calibration LED.

5. Begin testing according to standard test procedures. Allow 15 seconds of setting time after applying voltage before making accuracy measurements.

6. Under typical test conditions a minimum test time of 30 seconds is needed to reduce

test uncertainty to a level compatible with the accuracy of the Libra meter. (Check the instruction book for your test board or standard to determine the actual minimum test time.)

## 3.4 Service

The Libra meter is factory calibrated and requires no routine or scheduled service by the user.

## 3.5 Repair

Factory repair or replacement service is offered when you cannot fix a problem. If a problem can’t be corrected, return the whole meter to Wasion in the following paragraph

## 3.6 Returning a Meter

If you wish to return a meter, call your Wasion sales representative for a Return Authorization. The entire meter should be returned with the Wasion supplied Return

Authorization information form completed. Key information includes quantity, catalog number serial number(s) and a complete description of the problem. Your Wasion sales representative will provide return instructions.

## 3.7 Cleaning

## 3.8 Storage

The Libra Meter is a durable device; however, it should be handled and stored with care.

The temperature and humidity levels in storage are not critical; but extremes of either factor should be avoided

## 3.9 Troubleshooting Guide

Table 3-1 Troubleshooting

| No Display | - Circuit de-energized. - Test Link(s) Open. - Meter Internal wiring defective | - Rewire according to applicable diagram - Check that voltage and current connectors are seated properly. - Replace meter |
| --- | --- | --- |
| Meter runs slow | - Socket wiring error. - Meter Internal wiring defective - Defective sensor | - Rewire according to applicable diagram - Check that voltage and current connectors are sealed properly. - Replace meter |
| Meter overheats | - Meter socket has insufficient capacity or is not adequately wired - Meter is overloaded - Poor connection at socket terminal | - Replace meter socket with a heavy-duty model - Use transformer rated installation - Replace socket terminal |
| Error message on display | Meter detected an internal condition that may affect the meter data | Replace the meter and return to factory for service |

### 3.9.1 Self-Tests and Error Codes

The meter performs self-tests periodically. These self-tests are performed when the meter powers up and once every 24 hours thereafter. The current metering data is stored to non-volatile memory (NVRAM) prior to performing the periodic self-test. If the meter detects an error it will repeat the self-test every minute in order to attempt to recover from the condition causing the error. If the conditions causing the error clear, the meter will resume metering. If an error occurs, the meter will display “ER” in the primary display.

An error display indicates a fault in an electronic subsystem of the meter and indicates a condition that is not user serviceable.

# 4.0 Installation Instructions

## 4.1 General

CAUTION: Do not interchange power and meter module assemblies between meters.

Calibration data stored in meter is particular to a matched power and meter module. Interchanging these components causes the meter to lose calibration. The meter should never be disassembled. Failure to observe this practice can result in serious injury or death.

This chapter of the guide give instructions for the proper handling and installation of the Libra meter.

## 4.2 Inspection

Perform the following inspections when you receive the meter:

- Inspect for obvious damage to the cover, base, and meter assembly.

- Be sure the optical connector is free of debris.

- Compare the meter and register nameplates to the record card and invoice. Verify

the type, class, voltage, form number, and other pertinent data.

- Save the original packing materials.

## 4.3 Storage

Store the meter in a clean, dry (relative humidity < 50%) environment between -40℃ to

+85℃ (-40℉ to +185℉). Avoid prolonged storage (more than one year) at temperatures above +70℃ (+158℉). Store the meter in the original packing material.

## 4.4 Unpacking

As with all precision electronic instruments, the meter should be handled with care in

an outdoor environment. Follow these precautions when handling the meter:

- Avoid damaging the meter base, cover, reset mechanism (if supplied), and optical

connector (if supplied).

- When handling personality modules, grip the circuit board by its edges. Do not touch

the liquid crystal display

## 4.5 Selecting a site

The meter is designed and manufactured to be installed in an outdoor environment, at

operating temperature ranges between -20℃ to +70℃ (-40℉ to +185℉). Operation in moderate temperatures increases reliability and product life.

## 4.6 Meter installation

The meter is plugged into a meter socket using standard meter installation practices. The meter base has current and terminals extending outside from the back of the meter. These terminals engage with the socket jaws, which are connected to the service lines. The socket jaws provide heavy contact force with the help of spring. In some heavy-duty sockets, contact force is provided by a handle or wrench.

On powering-up the meter, verify meter operations by observing the display:

- For the first 3 seconds, the LCD display All Segment Test (all display items shown). If this all segment display is continued after 3 seconds then there is possibly an error in the installation wiring.

- LCD display the software version of the meter for 1 second.

- After that, LCD displays the number of digits for Energy.

- Verify the expected direction of energy flow on the display Annunciator.

# 5.0 Site Analysis Guides and Diagram

NOTICE:

These site analyses include basic connection diagrams for identification of metering installation.

Caution: these diagrams are not metering installation guides.

## 5.1 Form 1S (Self-Contained)

Figure 5.1 2-wire, 1-phase, 1-element

Caution: Actual installation procedures, equipment, and connections must conform to applicable codes and standards.

## 5.2 Form 2S (Self-Contained)

Figure 5.2 3-wire, 1-phase, 1-element

Caution: Actual installation procedures, equipment, and connections must conform to applicable codes and standards.

## 5.3 Form 3S (Transformer Rated)

Figure 5.3 2-wire, 1-phase, 1-element

Caution: Actual installation procedures, equipment, and connections must conform to applicable codes and standards.

## 5.4 Form 3S (Transformer Rated)

Figure 5.4 3-wire, 1-phase, 1-element

Caution: Actual installation procedures, equipment, and connections must conform to applicable codes and standards

## 5.5 Form 4S (Transformer Rated)

Figure 5.5 3-wire, 1-phase, 1-element

Caution: Actual installation procedures, equipment, and connections must conform to applicable codes and standards.

## 5.6 Diagrams

Figure 5.6 diagrams of ANSI C12.10 Internal Connection

# 6.0 Appendix

The Interface between Meter and AMR Module

Libra is equipped with two UART port, one is optical port, and the other is AMR port.

The AMR port meets to C12.18-1996 and C12.19-1997 standard. This port can be

used to communication with inner AMR module, such as RF/PLC/Zigbee etc.